[image: image1.jpg]DAKOTA ULTRASONICS

2 necessary
features to do the job.
They can go where you go,
to do the work you do,
saving you time and
money, making your
inspections fast and
accurate.

\ Ulira Bright Display
Broatdl & Narrow Bants

" Sguare Wave Pulser
(Variable & Fixed)

Digital Thickness
11008 Gain

DAC/TGC

Auto_Cal

Trig Fanction

Help Function

Large Memory,

20 Meter Range
Alphanumeric Notes
Rugged Aluminum Case
PC & Video Interface

SOUND SOLUTIONS





[image: image2.jpg]Physical
Si
9.3x5.3x7.9in. (237 x 135 x 200mm)
Weight:

9.41bs. (4.3Kgs) with Dry Cell Pack.
10.8Ibs. (4.9 Kgs) with NiCad Cells.

Case: Cast Aluminum Alloy to LM25 (BS1490)

Display: High brightness color TFT LCD
panel. Display area 3.8" x 3.03" (96 x 77mm)
320 x 256 pixels. A-Scan area 255 x 200
pixels for high quality signal display and
resolution. Removable screen for easy
cleaning. No parallax error.

Transducer Connectors: available with
BNC or “00" LEMO connectors.

Temperature:

Operating  ~10 to +55°C. 14 to 131°F
~20 to +70°C. -4 to 158°F
(survivable)

Storage 4010 +75°C. ~40 to 167°F

Environmental: Meets IP65 requirements.
Ultrasonic Specifications

System Linearity: Vertical = 1% Full Screen
Height (FSH) Amplifier Accuracy £0.1%dB.
Horizontal £0.4% Full Screen Width (FSW).

Pulser: 244 Fixed square wave pulse linked
to frequency selection. Tx width 40ns min,
250ns max. 344 Variable square wave pulser
from 40ns to 500ns duration. 200V peak ampli-
tude with rise/fall times <20ns into 500hms.

Delay: 244 Calibrated delay from 0-200 in.

in 0.05in. steps (0 to 5000mm in 1mm steps)
at steel velocity. 344 Calibrated delay from
0-400in. in 0.05in steps (0-10000mm in 1mm
steps) at steel velocity.

Gain: 0 to 110dB. Adjustable in 0.5, 2, 6, 14
and 200B steps. Direct access to gain control
atall times.

Damping: 344 Selectable between 50,
100,180 & 400 ohms.

P.R.F. 244 Adjustable to 150, 250, 500 and
1000Hz. 344 Adjustable to 35, 63, 150, 250,
500 & 1000Hz maximum.

Frequency Range: 244 4 Narrow Bands
centered at MHz, 2.25MHz, 5MHz, & 10MHz.
Wide Band 1.7MHz to 14MHz (-3dB).

344 5 Narrow Bands centered at 1MHz,
2.25MHz, 5MHz, 10 MHz, & 15MHz. Broad
Band at 2MHz to 22MHz(-6dB) and 1MHz to
35MHz (-200B).

Memory

Thickness Logging: Storage for 2000 thick-
ness readings configured into Block/Location/

Distributed by:

Number. Calibration settings stored with each
Block. Maximum number of Blocks is 14.
Unlimited Location/Number values, maximum
combination of 2000 readings. Readings may
be reviewed, edited and printed as required.

Panel Memory: 20 storage locations for
retaining calibrations.

A-Scan Memory: 244 Maximum of 50
‘waveforms, 344 Maximum of 100 waveforms—
stored with complete panel settings.
Waveforms may be recalled on display,

printed or transferred via RS232 serial port.

Notes: Alphanumeric labeling for panel
and A-Log allows the user to enter Notes for
storage with A-scans.

Measuring:
Units: English (in), Metric (mm), or Time (us).
Test Modes: Pulse-echo and transmit/receive.

Measurement Modes:
Mode 1 Signal Monitor.

Mode 2 Depth and amplitude of first signal
in gate.

Mode 3 Echo-to-echo distance measurement.

Mode 4 Trigonometric display of beam path,
surface distance and depth of indication.

Mode 5 T-Min mode for holding minimum
thickness reading. Resolution to 0.001in.
(0.01mm) for distance measurement, or 1%
FSH for amplitude measurement. Large display
of measurement at top of A-Scan display.
Measurement mode selectable between peak
and flank.

Note: All measurement functions also available
in unrectified RF mode.

Test Range: 244 0.2-400in. (5 to 10000mm)
at steel velocity. 344 0.05-800in. (1mm to
20000mm) at steel velocity. Variable in 1,2,5
sequence, or continuously in 0.05in. (1mm)
Increments. Also from 1 to 5000 (us).

Velocity: 244 0.0393 to 0.3937 in./micro-
second (1000m/s to 9999m/s) 344 0101 to
6299 in./microsecond (256mVs to 16000mVs)

continuously variable.

Probe Zero: 0 to 999,999 s, continuously
variable.

Features

Peak Memory: For echodynamic pattern
determination.

Gate Monitor: Two fully independent gates
for echo monitoring and thickness measure-
ment. Start and width adjustable over full
range of unit, amplitude variable from 0 to
100% FSH. Bar presentation. Positive or
negative triggering for each gate with visual
and audible alarms.

Y,

WLTR,,

& Ly

Q) ¢
)

S
S

SOUND SOLUTIONS

THE DFX 244 & 344 SPECIFICATIO

Gate Expansion: Expands range (o width
of Gate 1

DAC: DA curves may be entered and digital-
Iy drawn on the display. Reference, 648,
~12dB, ~14dB curves may be selected. DAC
curve selected acts as gate for alarm outputs
and height measurement in DAC +dB. DAC
parameters stored with Panel Memory. Curves
comply with ASME and JIS Codes.

Display Freeze: For capturing current
A-scan image.

TCG: Time Corrected Gain, also known as
Swept Gain. 40dB dynamic range greater than
300B per microsecond and up to 10 points.
may be used, setting all signals initially to
80% Full Screen Height.

Reject: 244 50% suppressive reject. LED
warning when selected. 344 80% linear reject.
LED warning when selected.

AWS-DGS: Built in Software for evaluation
of defect indications in accordance with AWS
standards or the AVG/DGS Method.

Auto_Cal: Provides Automatic calibration
from two echoes.

Rectification: Full wave, positive or negative
half wave and unrectified RF.

Help Key: On screen context-sensitive help.
Power Source:

Power: 10 ‘D’ size NiCad or alkaline cells.
8 hours+ duration from fully charged NiCad
cells. LED indication of low battery status.

A battery pack, which charges in 3 hours, is
available as an option or a direct power pack
for line power operation.

Charger: Universal line power input 85 to 260
volts A.C. Auto timed Charge. 3 hour fast
charge for special battery pack.

Outputs

Data: Full bi-directional RS232 serial
interface to transfer parameters, thickness
readings and waveform memories.

Analog: Proportional outputs programmable
to distance or amplitude of signal. Unrectified
RF signal 80mV peak to peak. Synchronisation
output for start of scan. Gate trigger output

on relay contact.

Video: Composite video, full PAL, or NTSC
compatibil
50Hz (PAL Mode).

Warranty
2 year limited

MKC KOREA

#1208, KICOX Venture Center, 188-5,
Kuro-dong, Kuro-ku, Seoul, Korea
TELEPHONE : 02.804.3600

FAX : 02.893.0498

WEB SITE : www.mkckorea.com

E-MAIL : ndt@mkckorea.com

- Update rate: 60Hz (NTSC Mode);

02-105725M




